## Slide 1: Solar, STELLAR and lunar photometry Lionel Doppler @IPC+FRC - Davos 20210929

## Slide 2: DWD/MOL-RAO
MOL4
Radiation
Processes
MOL1
Remote Sensing
MOL-RAO
Observatory Lindenberg
MOL2
Boundary Layer
MOL3
In-Situ Sounding
Lionel Doppler
Solar, STELLAR and lunar photometry
IPC/FRC – Davos - 20210929

## Slide 3: MOL-RAO: 52.2 N, 14.1 E, 120 m
Starphotometer
“Strahlungsmesszentrale”
(radiation lab)
Radiation platform
Wind house (historic)
RAMSES
(Raman Lidar)
Weather station
„Wetterwarte“
50 m tower
Balloon hall
Remote sensing station
(Cloud radar, Doppler Lidar, Windprofiler)
GM Falkenberg (e.g. 99 m tower)
Lionel Doppler
Solar, STELLAR and lunar photometry
IPC/FRC – Davos - 20210929

## Slide 4: Radiation platform
Lionel Doppler
Solar, STELLAR and lunar photometry
IPC/FRC – Davos - 20210929

## Slide 5: Basis of AOD photometry
Atmosphere’s entrance
I0,λ
AOD ???
Instrument:
Filterradiometer (ex: λ = 500 nm)
Beer-Lambert law at wavelength „λ“:
AOD = Aerosol Optical Depth
Aerosol layer
Iλ
Known
(measurement + spectral computation)
Known (V0, RIMO, SEMI)
Measured
Sun: „V0“, Langley-plot
Moon: „RIMO (ROLO)“, Moon Langley-plot + NASA
Star: „SEMI“: Star-catalog + 2 Star-Method
Lionel Doppler
Solar, STELLAR and lunar photometry
IPC/FRC – Davos - 20210929

## Slide 6
Stellar Photometry/Radiometry/Spectrometry
Lionel Doppler
Solar, STELLAR and lunar photometry
IPC/FRC – Davos - 20210929

## Slide 7
Starphotometers in the world
Only 8!
© Liviu Ivanescu, presentation SCILLA, Lindenberg, September 2020
Lionel Doppler
Solar, STELLAR and lunar photometry
IPC/FRC – Davos - 20210929

## Slide 8
Why only 8?
A dome
A telescope
An operator
Lionel Doppler
Solar, STELLAR and lunar photometry
IPC/FRC – Davos - 20210929

## Slide 9
Why a telescope?
Need to focus on stars = points less than 0.01°
Lionel Doppler
Solar, STELLAR and lunar photometry
IPC/FRC – Davos - 20210929

## Slide 10: Principles of the stellar photometry
Atmosphere’s entrance
I0_inst
AOD ???
Instrument:
Filter radiometer (ex: λ = 500 nm)
Aerosol layer
Iλ
-
+ +
Beer-Lambert equation !
Lionel Doppler
Solar, STELLAR and lunar photometry
IPC/FRC – Davos - 20210929

## Slide 11: What we need
Atmosphere’s entrance
I0_inst
AOD ???
Instrument:
Filter radiometer (ex: λ = 500 nm)
Aerosol layer
Iλ
A measurement of (Filter radiometer or spectrometer)
A knowledge of I0_inst (or “V0”) thanks to the calibration
V0 depends on: The star and the instrument: Name is SEMI for stellar phot.
SEMI = Star Extra-terrestrial Magnitude Instrumental
Lionel Doppler
Solar, STELLAR and lunar photometry
IPC/FRC – Davos - 20210929

## Slide 12
How to get I0_inst (V0)
FIRST OPTION: Langley-Plot
That requires:
Stable atmosphere: stable turbidity ( and )
  Elevated and dry site
  Or: Transportable system
Solution for ESCALIBURE (EXtinction CAmera and LumInance
BackgroUnd Register) = starphotometer of Granada
Lionel Doppler
Solar, STELLAR and lunar photometry
IPC/FRC – Davos - 20210929

## Slide 13
How to get SEMI I0_inst (V0)
OTHER OPTION:
Two stars method (TSM)
Used in Lindenberg, Eureka, Sherbrooke, Ny-Alesund
(SPST = not transportable system):
TSM
OSM
STP
U02
Lower Star
U2
m2
U0S
Only Star
US
m1
U01
Upper Star
U1
m1
U01, U02, U0S: stars‘ magnitude // U1, U2, US: Measured magnitude // m1, m2, mS: airmasses
→We need U01 – U02 SEMI1 – SEMI2 (Stars Extraterrestrial Magnitude Instrumental: I0_inst for astronomy) from upper star to lower star.
→ But U01 ~ ln(I0inst1) = ln(f◦I01) → U01 – U02 = ln(f◦I01/f◦I02) = ln(I01/I02) = M1-M2:
We only need Stars Extraterrestrial Magnitude (SEM or I0) -> Catalog
Beer-Lambert
Lionel Doppler
Solar, STELLAR and lunar photometry
IPC/FRC – Davos - 20210929

## Slide 14
How to get I0_inst (V0)
TSM
OSM
STP
U02
Lower Star
U2
m2
U0S
Only Star
US
m1
U01
Upper Star
U1
m1
U01, U02, U0S: stars‘ magnitude // U1, U2, US: Measured magnitude // m1, m2, mS: airmasses
→ With SEM Catalog and TSM (U1 and U2 meas.) we can make monitoring of optical depth and AOD and get rid of instrumental function (no need SEMI)
→ With optical depth and U1 and U2 measurements we get SEMI1 and SEMI2, we store these values. Later OSM = one star photometry with Star 1 and Star 2
Lionel Doppler
Solar, STELLAR and lunar photometry
IPC/FRC – Davos - 20210929

## Slide 15
Two stars measurements: summary
TSM
OSM
STP
U02
Lower Star
U2
m2
U0S
Only Star
US
m1
U01
Upper Star
U1
m1
U01, U02, U0S: stars‘ magnitude // U1, U2, US: Measured magnitude // m1, m2, mS: airmasses
→ TSM for having OD (then AOD) without “inst” part (without SEMI) of star magnitude but with the star magnitude independent of the instrument (SEMI)
→ TSM for calibration: compute SEMI1 and SEMI2 from TSM measurements for future OSM (one star method) photometry with star 1 and star 2
Lionel Doppler
Solar, STELLAR and lunar photometry
IPC/FRC – Davos - 20210929

## Slide 16
Limits of TSM method
TSM
OSM
STP
U02
Lower Star
U2
m2
U0S
Only Star
US
m1
U01
Upper Star
U1
m1
U01, U02, U0S: stars‘ magnitude // U1, U2, US: Measured magnitude // m1, m2, mS: airmasses
→ + We do not need a so stable atmosphere as for Langley-Plot
→ But we need homogeneous sky: OD same for lower and upper star
→ We need valuable SEM → good catalogue. Pulkovo Observatory catalog (Saint Petersburg) is our reference. But pretty old (stars change, resolution and precision standards of older time)
Lionel Doppler
Solar, STELLAR and lunar photometry
IPC/FRC – Davos - 20210929

## Slide 17
Stellar photometry targets
Zeta Draconis (ζ Dra) HR 6396 – Aldhibah
Beta Ursae Majoris (β UMa) HR 4295 – Merak
Eta Ursae Majoris (η UMa) HR 5191 - Alkaid
Beta Tauri (β Tauri) HR 1791 - Elnath
Kappa Orionis (κ Ori) HR 2004 - Saiph
Lionel Doppler
Solar, STELLAR and lunar photometry
IPC/FRC – Davos - 20210929

## Slide 18
Stellar photometry: summary
Only 8 active stations in the world
  Because infrastructure, maintenance and personal costs (dome, telescope, operator)
Calibration: Langley Plot or TSM (two star method)
Two options: Langley Plot + filter radiometer (ESCALIBURE, Granada)
or TSM and spectrometer (Dr. Schulz & Partner SPST: Canada, Ny-Å, Lindenberg)
Langley Plot = need portable, exclude low stars, need a lot of calibration (many stars)
TSM original, efficient, but dependency in SEM (Star magnitude) catalogue
TSM for AOD monitoring or TSM for calibration (SEMI) for 1 star method normal photometry with star 1 and star 2
TSM or Langley-Plot gives a set of +/- 8 stars for which we have SEMI an make then normal AOD photometry
Lionel Doppler
Solar, STELLAR and lunar photometry
IPC/FRC – Davos - 20210929

## Slide 19
Specificity and +/- of Lunar / Solar / Stellar
Star-/Lunarphotometry
Since 2010ies (2015)
Since 1990ies (1992)
Sun-photo-metry
Since 1980ies (1986)
Lionel Doppler
Solar, STELLAR and lunar photometry
IPC/FRC – Davos - 20210929

## Slide 20
Lunar or stellar?
Lionel Doppler
Solar, STELLAR and lunar photometry
IPC/FRC – Davos - 20210929

## Slide 21
Lunar or stellar?
SCILLA. Summer Campaign Solar+Lunar+Stellar, Lindenberg 2020
Lionel Doppler
Solar, STELLAR and lunar photometry
IPC/FRC – Davos - 20210929

## Slide 22: Photometry Instruments at MOL-RAO
PFR PMOD-WRC Sun GAW AOD (8)
CE318T Cimel Sun+Moon Aeronet AOD (6 o. 8), PWV
POM Prede Sun ESR(SkyRad) AOD (11)
SPST7 Schulz&Partner Star Ø AOD (17), PWV
PSR PMOD-WRC Sun + Glob Ø [315– 1020 nm]/1024
Lionel Doppler
Solar, STELLAR and lunar photometry
IPC/FRC – Davos - 20210929

## Slide 23
Solar = oldest and most certified and efficient technique
  Since the 80ies
  Portable and compact instrument
  Great signal/noise ratio
  Enough signal for UV, NIR, Ozone, (other techniques only VIS, AOD and WV)
  Possibility scattering measurements and extract aerosol properties (SSA, Phase function)
Lunar:
  Use same instrument as solar = compact, easy calibration
  Difficulty of reflectance model, but this is almost mastered
  Only Half of the cycle, only 2/3 of the night
Stellar:
  Expansive, take place, calibration difficult, need (partially) an operator
  Problem of star magnitude catalogue
  But we can choose the target and measure almost the whole night, heterogeneity is possible
Specificity Lunar/stellar/solar: summary
Lionel Doppler
Solar, STELLAR and lunar photometry
IPC/FRC – Davos - 20210929

## Slide 24
THANKS for your attention